## Slide 1: Al-Mustaqbal University. College of Engineering and Engineering Technologies. Biomedical Engineering Department.
Subject: Biomedical Instrumentation Lab.
Class (code): 5th
Lecture: 9
By: Zainab Sattar Jabbar M.SC. IN BME
BME

## Slide 2: X ray machine

## Slide 3: Production of x-rays
X-rays are produced by ‘shooting’ high energy electrons on a piece of metal (tungsten or molybdenum), called the ‘anode’. The electrons are absorbed in this metal and cause the creation of X-ray radiation.
You get high energy electrons by taking low energy electrons and accelerating these by means of a high electric voltage. This high voltage is created by the X-ray generator and applied between the anode (+), attracting the electrons, and the cathode (-).
99% of the energy released is given off as heat and only 1% is in the form of X-ray radiation! Getting rid of the heat is a major issue in tube design and operation.

## Slide 4: X-ray tube
External structures
Support structures
Protective housing
Glass or metal encloser
Internal structures
Anode and cathode

## Slide 5: X-ray tube support structures
Xray tube and housing assembly are quite heavy
Required support mechanism so radiologist can position them
Mainly ceiling, floor or C arm support system

## Slide 6: Protective Housing
When x-rays are produced, they are emitted isotropically
  That is, with equal intensity in all directions
Only x-rays emitted through window are called useful beam
  X-rays that escape through protective housing: leakage radiation
  Leakage radiation contributes nothing to diagnostic information and result in unnecessary exposure of patient and radiologic technologist
Protective housing guards against excessive radiation exposure and electric shock
  Also mechanically protects x-ray tube

## Slide 7: Metal or Glass Enclosure
X-ray tube is an electronic vacuum tube with components contained within a glass or metal enclosure
  vacuum allows for more efficient x-ray production and longer tube life
As glass enclosure tube ages, some tungsten vaporizes and coats the inside of glass enclosure
  Alter electrical properties of the tube, allowing tube current to stray and interact with the glass enclosure resulting in arcing and tube failure
  Most common cause of tube failure
Metal enclosures maintain constant electric potential between electrons of tube current and enclosure
  Longer life and less likely to fail
  Virtually all high-capacity x-ray tubes now use metal enclosures

## Slide 8: X-ray tube
X-ray Out
High Voltage In
X-ray Tube ‘Insert’; the inside is vacuum
X-ray Tube Container
protects the tube from mechanical shock and absorbs stray X-ray radiation
Anode disc rotation mechanism (‘rotor’)
anode disk
kathode

## Slide 9: X-ray tube
©
In high end X-ray tubes, the anode is rotating, so that it can absorb more heat. Temperatures of 3000oC can be reached.
Also, the anode may be cooled with water or with oil, that is circulated through the tube and cooled with a ‘heat exchanger’ outside the tube.
tubes with cooling fluid
The place on the anode where X-rays are generated is called the ‘focal spot’.

## Slide 10: Cathode
Cathode is the negative side of the x-ray tube
It has two primary parts, a filament and a focusing cup
  Dual-filament cathode allows two focal spots (e.g., 0.5 and 1.5 mm)
  Focusing cup is a metal shroud that surrounds filament
  Tube current is adjusted by controlling filament current

## Slide 11: Anode
Anode is the positive side of the x-ray tube
Two types: stationary (dental) and rotating (general purpose)
  Higher tube currents and shorter exposure times are possible with rotating anode because of their better heat dissipation
Three functions in anode of the x ray tube:
  Electrical conductor that receives electrons emitted by cathode and conducts them through the tube to the connecting cables and back to the high-voltage generator
  Mechanical support for the target
  Thermal dissipation

## Slide 12: Target
The target is area of anode struck by electrons from cathode

## Slide 13: Focal Spot
Focal spot is the area of target from which x-rays are emitted
  The smaller the focal spot, the better the spatial resolution of the image
  Unfortunately, as the size of focal spot decreases, heating of target is concentrated onto a smaller area (limiting factor to focal spot size(
Line-focus principle: angling target makes effective area of the target much smaller than actual area of electron interaction

## Slide 14: X-ray Image Quality
Most user complaints on a functioning X-ray systems are about reduced Image Quality. The factors that determine image quality are:
  spatial resolution or sharpness
  image contrast
  image noise
  image artefacts
Each of these factors can lead to reduced visibility of image details and to a sub-optimal/wrong diagnosis.
Some image quality reduction occurs in case of high X-ray absorption, as sometimes caused by obese patients.
This can lead to low contrast and high noise in the image. It can sometimes not be improved by improving the system.
Chest X-ray of normal (left) and obese (right) patient

## Slide 15
THANK YOU