AL- Mustaqbal University

Optometry Department

Lec.12

Optical Instruments

assistant teacher

MCS

Ali Hadi Al-Husseini

2023-2024

Indirect ophthalomscope

Definition:

is an necessary tool to diagnose and manage vitreoretinal disorders. Help observation of peripheral retina, it laid the basic foundation for the modern treatment of vitreo retinal diseases.

PRINCIPLE:

➔ The technique is called Indirect because the fundus is seen through a summarizing lens.

➔ The principle of indirect ophthalmoscopy is to make the observer's eye myopic by placing a strong convex lens (+13D to +30D) in front of the eye that real, reversed and laterally reversed image is formed in front of lens.

➔ When both the illuminating and reflected beam pass through the pupil, that area of fundus is seen.

➔ A pencil of rays can be drawn from the patient's fundus to the viewer's retina.

➔ An middle, reversed but real image of the patient's fundus is formed in the focal plane of the ophthalmoscopy lens, between the lens and the viewer.

➔ The observer must accommodate on this image.

➔ Normal interpupillary distance is about 60 mm

➔ Reduced to 15 mm by use of prism in IDO

FIELD OF VIEW

➔ The field of view is determined by the ratio of lens diameter and focal length. This ratio can be written as: Field of view = Lens diameter / Focal length

= Lens diameter x dioptric power

➔ Given lenses of equal power, a larger lens provides a wider field of view.

➔ If lenses have equal diameters, a stronger lens provides a wider field of view; however, after an optimum power, stronger lenses [more convex] tend to have smaller diameter, hence a stronger lens does not always provide a larger field.

MAGNIFICATION:

➔ Total magnification depends on lens power [inversely proportional] and observation distance.

➔ The stronger lens, the less magnified of the image but wider field of view.

The power of the condensing lens determines:

✓ Field of view

✓ Magnification

✓ Working distance

COMPENSATION FOR REFRACTIVE ERROR

➔ With the indirect method, minor changes in the observer's accommodation can compensate for major changes in the patient's refractive error.

➔ In indirect ophthalmoscopy, compensation for refractive error in patient's eye can be made without additional lenses.

➔ The presbyopic observer, who cannot change accommodation, can compensate by changing observation distance.

➔ Distance between observer 'O' and lens 'L' = 45 cm.

➔ Emmetropic eye: Aerial image at E' and observer need to accommodate upto 2.5D [40cm].

➔ Myopic eye [-5D]: Aerial image will be at M' and accommodation of 2.4D [41cm] is required to see it clearly.

➔ Hypermetropic eye [+5D]: Aerial image will be at H' and observer will have to accommodate upto 2.6D[38.4cm].

➔ Change of +/ - 0.1 D in observer's accomodation accounts for difference of +/ - 5 D in patient’s refractive error.

➔ Thus, minor changes in the examiner's accommodation can easily compensate for major refractive errors that patient may have.

➔ The presbyopic observer, who cannot alter accommodation, can compensate for the patient's refractive error by changing the observation distance or by using a near-vision aid.

Technique of Indirect ophthalmoscopy

➢ Head set adjustment

➢ Eye piece adjustment

➢ Light beam adjustment

➢ Choosing, positioning and technique to hold condensing lens.

➢ Positioning of the patient

➢ Scleral indentation

➢ The fundus drawing

Filters

- A. Green light - Nerve fibre layer, Blood vessels, microaneurysms

- B. Red light - Subtle pigmentary abnormalities

- C. Blue light - Angioscopy

- D. Yellow filter - Reduces photophobia Condensing Lens

Three types:

1. Biconvex

2. Plano convex

3. Aspheric

PREREQUISITES

- i. Indirect ophthalmoscope.

- ii. Dark room.

- iii. Convex lens.

- iv. Pupil of the patient should dilate.

Technique:

➔ Ideal position of the ophthalmoscope.

➔ Axis perpendicular to the visual axis of examiner.

➔ The scope not resting on the nose of the examiner.

➔ The eyepiece as close to the examiners pupils as possible.

➔ Adequately adjusted IPD.

For appropriate Examination:

➔ Proprioception helps patient to look in cardinal gazes.

➔ Patient should be urged to keep the other eye open.

Indirect ophthalmoscopy should be used:

✓ When examining a patient with symptoms of shadows, hazy vision, flashes of light, and floaters, which may suggest retinal tear, detachment or other retinal abnormality.

✓ When following a patients with systemic diseases that have retinal manifestations, such as diabetes, high blood pressure, cardiovascular disease, sickle cell disease, infectious and autoimmune diseases.

✓ In patient with cataract or following cataract extraction.

✓ For older patients who are at greater risk of developing glaucoma and age-related macular degeneration.

ADVANTAGES

1. Optimal working distance = 40 to 50 cms

2. Good illumination & stereopsis: better view in presence of media opacities

3. Ease of use with scleral indenter

4. Lenses from 13 to 30 D range

5. Image not affected by the patient’s refractive power

6. In children, eyes with nystagmus

7. Delivery of LASER, operating room for cryo/scleral buckling.

DISADVANTAGES

1. Inverted and inversed image

2. Low magnification

3. Dilation required

4. Difficult to master

THANK YOU