AL- Mustaqbal University

Optometry Department

Lec.11

Optical Instruments

assistant teacher

MCS

Ali Hadi Al-Husseini

2023-2024

Lensmeter

It is an instrument for determining the lens's power or degree in contact lenses or eyeglass. It exists by a number of names, including vertometre and focimeter. The image below illustrates the two types: manual and automatic. Utilized by opticians and optometrists.

Uses of lenemeter

➢ Single Vision lenses

➢ Bifocal lenses

➢ Trifocal lenses

➢ Progressive Addition lenses

➢ Determine the Sphere power

➢ Determine the Cylindrical lens power

➢ Mark the optical center

➢ Contact lens power

➢ Power of prism

Types lensmeter

1- manual lensmerer can be defined as an instrument used by an optometrist to find the back-vertex or the front vertex power of the eyeglasses. A manual lensometer is portable and can be carried anywhere. But a person needs to have a better idea to measure the power of a lens.

2- Autometed lensmeter It is a fully automatic well-programmed device primarily used in clinics. It is easy and faster and can print prescriptions. It is less accurate when compared to a manual lensometer.

Parts Of Manual Lensmeter:

✓ Eyepiece: It is that part of the lensometer through which the viewer can view the internal components of the lensmeter.

✓ Lens holder: It is used to hold the lens in its proper position.

✓ Lens table: It provides a resting place for the spectacle frame while neutralizing the power of the lens.

✓ table lever: It helps raise or lower the level of the spectacle table according to the user

✓ Lens marker: The handle control’s the pins and is used to mark the optical center or the prism reference points.

✓ Power wheel: It is used to clear the filter and find out the spherical and cylindrical power of the given lens.

✓ axis wheel: It rotates the target to find the cylindrical axis.

How to use Lensometer

1. The device is worked and the standard is set to find the clear target at

ZERO degrees by rotating the eyepiece clockwise and counterclockwise.

2. Put lens or the eyeglasses to be measured on lens table and lowering the

lens holder to fix in the center of the grid, and one of the hands must be

holding the power wheel.

3. Because of the placement of the lens in the device, the target in grid will

become unclear, so we move the power wheel up or down to obtain

accurate clarity.

4. If the target (the circle in green) is clear for both axes together, then the lens type is SPHERICAL

5. If the target is clear on one axis and the other is not clear, then this means that the lens is CYLINDRICAL.

Analyzing the Spherical lenses

▪ Move the lens such that the target is exactly at the center.

▪ Rotate the power drum until target is clear and sharp (all the dots at the center should be separate & sharp)

▪ Stop rotating the power drum when the target is sharp for the first time.

▪ Always count the upper mark on the power scale for accuracy of the spherical power of the lens.

Analyzing the Cylindrical lens

▪ The process of Neutralizing a toric lens when compared to spherical lens is completely a different concept.

▪ with the help of few tips we will be able to measure the lens power very fast and accurately.

▪ The key and the most important thing to remember is the central orientation of the dots.

➔ Step 1:-

Move the lens so that the target is at the center of the protractor. This gives the optical center location.

➔ Step 2:-

• Rotate the power drum so that the central lines (oriented in one direction).

➔ Step 3 :

Rotate the axis wheel so that one meridian of the target is parallel to the orientation of the central lines.

Note that only one meridian of the target will be sharp. As the 180° meridian is sharp in this picture, the target is clearly in the form of a cylinder (close and parallel horizontal lines), then we read the degree that represents the FIRST reading, as in the figure below.

Rotate the power drum to obtain the SECOND reading, which is perpendicular to the axis of the first reading, which will appear in the form of a cylinder (close and parallel vertical lines) as shown in the figure below

the first reading is the spherical lens, while the second reading is the cylindrical lens, and the axis is taken from the second reading

Measurement of Bifocal

For measuring the distance correction of a bifocal lens, the lens measured at the optical center. The optical center of a conventional bifocal lens is just above the center of the horizontal line of the bifocal lens and read the near correction from the center of the lower segment. For determining add power, turn glasses around to read from front vertex.

Measurement of Trifocal

As with the bifocal, read the distance correction just above the segment line, and read the near correction from the center of the lower segment.

Measurement of Progressive lenses

➢ Locate the center to make the reading accurate, otherwise is similar to the simple lensometery.

➢ All new pair contain sticker on it, which describe geometry of lens.

➢ The add power is read through the lower/nasal area of the lens

➢ Conventional progressive addition is measured from front vertex

➢ Automated is more suitable for this purpose.

Measurement of prism

➢ Prism moves (deflects) the lensometer target away from the center of the reticle.

➢ The target is deflected in the direction of the prism base.

➢ When viewing the mires, lines may be off-center.

➢ The rings inside the lensometer are measured at 1 prism diopter.

➢ The number of rings from the center of the reticle that the lines are over is the amount of prism in the glasses.

Limitations of Manual focimeter

✓ Although it is good for single vision lenses, but may misguide in addition determination of bifocal/trifocal lenses.

✓ Power varies with 0.25D difference don't calculate 0.12D difference. O Difficult in case of progressive lenses.

✓ Can cause damage to the contact lenses.

Automated Focimeter

Principle:

The degree to which a beam of light is deflected as it passes through the lens depends on the focal and prismatic power of the lens and the distance from its optical center.

Working

1. A square pattern of four parallel beam of light is passed through the lens to be tested.

2. It is quick and easy, minimum expertise required to handle.

3. Accuracy is variable with machine to machine.

Important Uses:

It can used for accuracy in all conditions:

➢ Proper centration/ decentration

➢ Axis Marking.

➢ Measuring power of near addition lens (bifocal, trifocal, progressive) difference.

➢ IPD measurement of lens mounted spectacle.

THANK YOU