## Slide 1: A practical lecture of Capnography

## Slide 7: 0

## Slide 8
Mainstream | Sidestream
CO2 sensor is housed external to the breathing circuit | CO2 sensor is inserted between the breathing circuit and endotracheal tube
No gas removed from circuit | Gas is constantly aspirated from circuit via a 6 feet sampling tube into the unit containing CO2 sensor
ncrease in mechanical dead space | Minimal dead space
Waveform in real time | Waveform is delayed (1–4 seconds) due to transportation of gases from the patient’s airway to the unit containing the sensor

## Slide 11
Time capnogram and its components.
Time capnogram showing segments, phases and angles.
The inspiratory segment is phase 0
The expiratory segment is divided into three phases: I, II, and III.
The maximum value of CO2 at the end of the breath is designated as end-tidal PCO2 (PetCO2). It is lower than arterial CO2 by about 5 mm Hg

## Slide 12: Prolonged phase II and steeper phase III suggestive of bronchospasm, or airway obstruction

## Slide 13: B. Expiratory valve malfunction resulting in elevation of the baseline This is due to rebreathing of expiratory gases from the expiratory limb during inspiration

## Slide 14: Inspiratory valve malfunction resulting in rebreathing of expired gases from inspiratory limb during inspiration

## Slide 15: Capnogram with normal phase II but with an increased slope of phase III. Observed in pregnant subjects under general anesthesia.

## Slide 16: Patient is attempting to breathe during partial muscle paralysis. Surgical movements on the chest and abdomen can also result

## Slide 17: The baseline is elevated as a result of CO2 rebreathing

## Slide 18: Esophageal intubation resulting in the gastric washout of residual CO2 and subsequent CO2 will be zero;

## Slide 19: Spontaneously breathing capnograms where phase III is not well delineated

## Slide 20: Dual capnogram in one lung transplantation patient. The first of peak of phase III is from the transplanted normal lung, whereas the second peak is from the native disease lung. A leak around the sidestream sensor port at the monitor can also result in a dual peaked capnogram. This is because of the dilution of expired PCO2 with atmospheric

## Slide 21: Malignant hyperpyrexia where CO2 is raising gradually with zero baseline suggesting increased CO2 production with CO2 absorption by the soda lime

## Slide 22: cardiogenic oscillations

## Slide 23: Sudden raise of baseline and the partial pressure of end-tidal carbon dioxide (PetCO2) due to contamination of the sensor with secretions or water vapor

## Slide 24: Intermittent mechanical ventilation (IMV) breaths in the midst of spontaneously breathing patient

## Slide 25: Cardiopulmonary resuscitation: Capnogram showing positive waveforms during each compression suggesting effective cardiac compression generating pulmonary blood

## Slide 26: increase in partial pressure of end-tidal carbon dioxide (PetCO2) values suggest spontaneous return of circulation during cardiopulmonary resuscitation

## Slide 27: Capnogram showing rebreathing during inspiration. This is normal in rebreathing circuits such as Mapelson D or Bain circuit

## Slide 28: Home work

## Slide 29: Home work

## Slide 30: Home work

## Slide 31: Home work

## Slide 32: Home work

## Slide 33
Thank you